Instructions to User

Dear Users, thank you very much for purchasing our product. Please read the manual very carefully before using this device. Failure to follow these instructions can cause measuring abnormality or equipment damage.

All rights reserved.

Notes:

- The contents contained in this manual are subjected to change without notice.

- Information furnished by our company is believed to be accurate and reliable. However, no responsibility is assumed by us for its use, or any infringements of patents or other rights of third parties that may result from its use.

- This Manual is written and compiled in accordance with the council directive MDD93/42/EEC for medical devices and harmonized standards. In case of modifications and software upgrades, the information contained in this document is subject to change without notice.

- The Manual describes, in accordance with the Pulse Oximeter’s features and requirements, main structure, functions, specifications, correct methods for transportation, installation, usage, operation, repair, maintenance and storage, etc. as well as the safety procedures to protect both the user and equipment. Refer to the respective chapters for details.

Instructions for Safe Operations

  - Check the device to make sure that there is no visible damage that may affect user’s safety and measurement performance about sensors and clips. It is recommended that the device should be inspected once a week at least. When there is obvious damage, stop using the device.

  - Necessary maintenance must be performed by qualified service personnel ONLY. Users are not to service it by themselves.

  - The oximeter can not be used together with the devices not specified in User Manual.

Cautions

  - Explosive hazard—DO NOT use the oximeter in environment with inflammable gas such as some ignitable anesthetic agents.

  - DO NOT use the oximeter while the testee is under MRI or CT scanning.

- The uncomfortable or painful feeling may appear if using the sensor of this device continuously on the same place for a long time, especially for the poor microcirculation patients. It is recommended that the sensor should not be applied to the same location for over 2 hours. If any abnormal condition is found, please change the position of sensor.

- For each individual, we suggest consulting a physician for best proper measurement point. The device can not be clipped on the edema and tender tissue.

- The light (the infrared light is invisible) emitted from the device is harmful to the eyes, do not stare at it.

  - To dispose the device, the local law must be followed.

Attentions

    - Keep the oximeter away from dust, vibration, corrosive substances, explosive materials, high temperature and moisture.

    - The device should be kept out of the reach of children.

    - If the oximeter gets wet, please stop operating it. When it is carried from cold environment to warm and humid environment, please do not use it immediately.

    - DO NOT operate the button on front panel with sharp materials.

    - High temperature or high pressure steam disinfection to the oximeter is not permitted. Refer to related Chapter for instructions of cleaning and disinfection.

Intended use

    - For spot/continuous check of blood oxygen saturation level at home, in sport and recreational use, such as mountain climbing, high-altitude activities and running.

    - Not intended for self medical diagnostic use or to replace a doctor's visit, please contact your physician for any measurement readings that you feel uncomfortable with.

Contents

1 Overview	2

1.1 Features	2

1.2 Applications and scope	2

1.3 Environment requirements	2

2 Principle	2

3 Battery installation	2

4 Technical specifications	2

4.1 Main functions	2

4.2 Main Parameters	2

5 Installation	2

5.1 View of the front panel	2

5.2 Wristband installation and probe connection	2

5.3 Accessories	2

6 Operating Guide	2

6.1 Application method	2

6.2 Change display direction	2

6.3 Enter and exit the clock interface	2

6.4 Pause alarm	2

6.5 Menu operations	2

6.6 Clock setting	2

6.7 Real-time data transmission setting	2

6.8 Data storage setting	2

6.9 Uploading the data to the PC after recording	2

6.10 Power off	2

6.11 Exit the main menu	2

6.12 Charging the battery	2

6.12   Attention for operation	2

6.13 Clinical restrictions	2

7 Maintenance、transportation and storage	2

7.1 Cleaning and Disinfecting	2

7.2 Maintenance	2

7.3 Transportation and storage	2

8 Troubleshooting	2

9 Key of Symbols	2

	Function Specification	2

Appendix 1	2

# 1 Overview

The pulse oxygen saturation is the percentage of HbO2 in the total Hb in the blood, so-called the O2 concentration in the blood. It is an important bio-parameter for the respiration. A number of diseases relating to respiratory system may cause the decrease of SpO2 in the blood, furthermore, some other causes such as the malfunction of human body's self-adjustment, damages during surgery, and the injuries caused by some medical checkup would also lead to the difficulty of oxygen supply in human body, and the corresponding symptoms would appear as a consequence, such as vertigo, impotence, vomit etc. Serious symptoms might bring danger to human's life. Therefore, prompt information of patients' SpO2 is of great help for the doctor to discover the potential danger, and is of great importance in the clinical medical field.

The Pulse Oximeter features in small volume, low power consumption, convenient operation and being portable. It is only necessary for patients to put one of his fingers into a probe for diagnosis, and a display screen will directly show the measured value of pulse oxygen saturation with the high accuracy and repetition.

## 1.1 Features

- The device can accurately measure SpO2.

- The device will power off automatically if no signal is detected within 5 seconds on the measuring interface.

- Rechargeable battery is used for operation, and the continuous working time is 30 hours.

- When Low-Voltage Warning is shown on screen, the device maybe not operate normally.

- Light and portable, total weight with batteries is about 52g.

- Classification: Class II a, （MDD93/42/EEC  IX  Rule 10）

## 1.2 Applications and scope

The Pulse Oximeter can be used in measuring the pulse oxygen saturation and pulse rate through finger. The product is suitable for being used in family, hospital, oxygen bar, community healthcare, sports, high altitude activities, mountain climbing (It can be used before or after doing sports, and it is not recommended to use the device in the middle of sport activities) and etc. This device is not intended for medical diagnostic use or to replace a doctor's visit and self diagnostics. Please consult your physician for any measurement readings that you feel uncomfortable with.

The problem of overrating would emerge when the patient is suffering from toxicosis which caused by carbon monoxide, the device is not recommended to be used under this circumstance.

## 1.3 Environment requirements

Storage Environment

- Temperature :-40°F ~140°F (-40℃～+60℃)

- Relative humidity :5%～95%

- Atmospheric pressure :500hPa～1060hPa

Operating Environment

- Temperature: 50°F ~104°F (10℃～40℃)

- Relative Humidity :30%～75%

- Atmospheric pressure:700hPa～1060hPa

# 2 Principle

Principle of the Oximeter is as follows: An experience formula of data process is established taking use of Lambert Beer Law according to Spectrum Absorption Characteristics of Reductive Hemoglobin (Hb) and Oxyhemoglobin (HbO2) in glow & near-infrared zones. Operation principle of the device is: Photoelectric Oxyhemoglobin Inspection Technology is adopted in accordance with Capacity Pulse Scanning & Recording Technology, so that two beams of different wavelength of lights can be focused onto human nail tip through perspective clamp finger-type sensor. Then measured signal can be obtained by a photosensitive element, information acquired through which will be shown on screen through treatment in electronic circuits and microprocessor.

Figure 1.

# 3 Battery installation

- No battery needed to install since it come with rechargeable battery

# 4 Technical specifications

## 4.1 Main functions

A. SpO2 value display

B. Pulse rate value display, bar graph display

C. Pulse waveform display

D. Low-voltage indication: low-voltage indicator appears before working abnormally which is due to low-voltage

E. The display mode can be changed

F. Screen brightness can be changed

G. A pulse sound indication

H. With alarm function

I. With SpO2 value and pulse rate value of storage, the stored data can be uploaded to computers

J. It can be connected with an external oximeter probe

K. Real-time data can be transmitted to computers

L. With clock function

## 4.2 Main Parameters

A.  Measurement of SpO2

Measuring range: 0%～100％

Accuracy:

When the SpO2 measuring range is 70%～100％,the permission of absolute error is ±2％；below 70% unspecified

B.  Measurement of pulse rate

Measuring range: 30bpm～250bpm

Accuracy: ±2 bpm or ±2% (select larger)

C.  Resolution

SpO2: 1%, Pulse rate: 1bpm.

D. Measurement Performance in Weak Filling Condition

SpO2 and pulse rate can be shown correctly when pulse-filling ratio is 0.4%. SpO2 error is ±4%, pulse rate error is ±2 bpm or ±2% (select larger).

E.  Resistance to surrounding light

The deviation between the value measured in the condition of man-made light or indoor natural light and that of darkroom is less than ±1%.

F.  Power supply requirement: : 3.6 V DC ~ 4.2V DC.

G.  Optical Sensor

Red light (wavelength is 660nm，6.65mW)

Infrared  (wavelength is 880nm,  6.75mW)

H.  Adjustable alarm range:

SpO2 ： 0%~100%

Pulse Rate: 0bpm~254bpm

# 5 Installation

## 5.1 View of the front panel

Figure 2. Display description

Figure 3. left and right view

- Probe jack: It is used to connect a SpO2 sensor to measure the oxygen saturation and pulse rate.

- USB port: It is used to connect a personal computer to upload stored data or charge the lithium battery via a data line.

- Charging indication light. When the device is in the state of charging, the indication light is orange, and when the battery is fully charged, the light turns to green.

## 5.2 Wristband installation and probe connection

A.  Put one side of the wristband on which there is no ring through the belt hole above the device, then put it through the other belt hole under the device.

B.  Make the alignment signs on the probe in accordance with the alignment signs on the device, then plug the probe into the jack.

C.  It appears as the following figure after installation.

Figure 4. Front view

## 5.3 Accessories

A. Wristband

B. User Manual

C. Power adapter/charger

D. USB data line

E. Software installation disk (PC software)

F. An adult probe（Model:CMS04）An infant probe is optional and may purchase seperately

G..Optional wireless connection kits

# 6 Operating Guide

## 6.1 Application method

A.  Install the wristband according to the instructions of chapter 5.2.

B.  Put the suitable probe into the jack on the right side of the oximeter. (Use the probe that is come with the package or purchased through our distributor; never replace it with the similar ones by other manufacturers).

C.  Put the finger into the probe.

D.  Turn on the device by long pressing the button on the panel.

E.  Do not shake the finger and keep the patient in a stable state during the process.

F.  The data can be read directly from the screen.

Fingernails and the luminescent tube should be on the same side.

If the alarm is on, the device will provide medium-priority alarm signal when probe or finger is out. Intermittent alarm will occur and the user interface presents "FINGER OUT".

Medium priority indicating that prompt operator response is required.

Figure 5. In operation

(Actual probe may be different with the probe as figure 5）

Operation terms:

CLICK: short press of power button

PRESS: press and hold power button (1sec)

## 6.2 Change display direction

On the measuring interface, enter the clock interface by clicking the button, and then click the button to change the display direction within 30 seconds.

## 6.3 Enter and exit the clock interface

a)  On the measuring interface, click the button in order to enter the clock interface, and it will automatically return to the measuring interface if there is no more operations within 30 seconds.

b)  On the measuring interface, press the button for about 10 seconds in order to enter the clock interface, and the device would return to the measuring interface by Press the button for about 10 seconds again.

## 6.4 Pause alarm

a)  Alarm including the alarm of measure data's going beyond the limits, the alarm of low-voltage, the alarm of probe or finger's out of position.

b)  On the measuring state, if the alarm function is turned on and an alarm can be paused by clicking the button, but the function will be renewed in about 60 seconds.

c)  If you want to turn off the alarm, you should enter the menu for that operation.

## 6.5 Menu operations

When the device is under the measuring interface, Press the button for about 1 second in order to enter the menu interface shown as figure 6. Users can adjust the setting through the main menu, such as backlight, alarm, clock, data transmission (with the data line), data storage, and power off. The specific operation methods are as the following:

Figure 6. Main Menu Interface

a)  Backlight adjustment

On the main menu interface, click the button to select “Brightness”, Press the power button and hold to adjust the backlight brightness.

b) Alarm setting

On the main menu interface, click the button to select “Alarm”, Press the power button (1sec) to enter the alarm setting interface as shown in Figure 6:

1) Adjusting the high and low limits of alarms

Click the button to select “Direction”, then press the button to choose Up or Down. (this will be the direction the value of the high-low limits of SpO2 and pulse rate will be adjusted)

To raise the SpO2 and pulse rate limit, choose “Direction” as ‘up’, then Click the button to highlight the parameter to be adjusted: SpO2 high limit (SpO2 ALM HI), SpO2 low limit (SpO2 ALM LO), Pulse rate high limit (PR ALM HI), Pulse rate low limit (PR ALM LO). Press the button and hold to adjust the selected limit to the desired higher value and release the button once the higher limit has been reached.

To lower the SpO2 and pulse rate limit, choose “Direction” as ‘down’, then click the button to choose the parameter to be adjusted. Press the button and hold to adjust the selected limit to the desired lower value and release the button once the lower limit has been reached.

If the alarm function is on, the device will provide medium-priority alarm signal when the data of SpO2 or pulse rate is beyond the limit. Intermittent alarm will occur and the measurement shows in yellow.

Medium priority indicating that prompt operator response is required.

2) Turn alarm on/off

Click the button to select “Alarm”, press the button to choose alarm on or off, press “on” to turn on the alarms and “off” to turn off the alarms.

Figure 7.Alarm Setting Menu

3) Pulse sound indication setting

Click the button to select “Pulse Sound”, Press the button to choose to have the Pulse Sound (heart beat) alarm “on” or “off”.

4) Exit the Alarm settings

Click the button to select “EXIT”, Press the button to exit the Alarm Settings Menu.

## 6.6 Clock setting

On the main menu interface, click the button to select “Clock” to enter the clock setting interface by press the button.

Figure 8. Clock Setting Menu

a) When entering the clock setting menu, the menu choice bar would be on the item of “set time”, and the state would always be “no” whenever it enters the clock setting menu on the purpose of avoiding unexpected changes of time due to improper operation. You can change the state by press the button, choose “yes” to reset the time, choose “no” to forbid time resetting.

b) Click the button to select the parameter that you want to change, then adjust the data by press the button.

c) Click the button to select “Exit”, exit the clock setting menu by press the button. If you have reset the time or date, when exiting the clock setting menu, firstly the renewed time and date would be displayed on the screen, then it returns to the main menu; if you did’t reset the time and date , when exiting the clock setting menu, the device would return to the main menu directly.

## 6.7 Real-time data transmission setting

First install the software into the computer, and two icons would appear on the desktop after installation. The icon of "SpO2 "is a program for receiving real-time data shown as figure 9; the icon of "SpO2 Review" is a program to review the stored data which is shown as figure 10.

a)　Please connect the device to computer with the USB data line, then double click the "SpO2 "icon to start the program.

b)　On the main menu interface, click the button to select “Usb”, press the button to choose whether transmit the real-time data to computer which displays the data synchronously or not, choose “on” to permit transmission, choose “off” to forbid transmission

c)　When you unplug the data line from computer, there is a dialog box "Save data at view" appearing on the desktop, in which you can input some patient's basic information.

Figure 9. SpO2 program

Figure 10. SpO2 Review program

If the users choose to turn on the synchronizing display function on computer, it would probably take several seconds for the data to appear on the computer screen.

## 6.8 Data storage setting

This model has the ability to store 24 hours’ measurement data. It can store the measured pulse rate and SpO2 value accurately, transfer the data to the computer, display the data and print reports (with the included SpO2 Software - Green Heart)

a)  On the main menu interface, click the button to select “Record”, then Press the button to choose whether store the data or not, choose “Yes” to permit storing, choose “No” to forbid storing.

b) The device can store automatically.

c) If the data storage function is being turned on, when return to the measuring interface, a red "REC" sign and a flashing red dot would appear on screen, which means the device is in a state of storing.

d) In the state of storing, no matter what the interface the device is on (measuring interface, menu interface, clock interface), the sign "Recording" would appear on the screen in 30 seconds, then the clock interface would appear after several seconds. then the screen will be automatically shut down. Click the button at this moment, the sign "Recording" would appear on the screen, and then the screen will be automatically shut down again; if press the button, the device would return to the former interface.

e) If turning on the data storage function, the former data storage will be automatically removed.

f) In the state of data storing, after the screen is automatically shut down, the pulse sound indication would be off for saving power.

g) When the storage space is full, it displays “Memory is full” on the screen, and then shut down in a few seconds. But it will still display “Memory is full” by the next time you turn on the device on the purpose of warning the user, if click the button again, it will enter the measuring interface.

## 6.9 Uploading the data to the PC after recording

a)　Please connect the device with computer by the data line which is equipped with the device, then double click "SpO2 Review" icon to open"SpO2 Review" program, click the ‘New Session’ Icon in the software, enter the patient data and then click ‘ok’.  The Software will then display “device connected, waiting for data”.

b)　At this time, press the button to enter the “Main Menu” and then Click the button to select “Upload”. click the button to select “on” then the data will be transferred to your computer.

c)　In the state of storing, it is not applicable for the users to upload the stored date to computer.

d) When the upload of stored data is finished, the menu choice bar will move to "Exit" automatically.

## 6.10 Power off

On the main menu interface, Click the button to select "Power off", then Press the button to shut down the device.

## 6.11 Exit the main menu

On the main menu interface, Click the button to select "Exit", then Press the button to exit the main menu.

## 6.12 Charging the battery

There are two kinds of charging methods:

a） Connect the device with computer by data line, then the device should be in charging state.

b)  Plug in the supplied power adaptor, the device will be in charging state.

c)  When the device is in the charging state, the indication light appears to be orange, and when the battery status is full, the light turns to green.

If the alarm function is on, the device will provide high-priority alarm signal when the battery is in low power status .Intermittent alarm will occur and the battery icon turns red in the state of flashing.

High priority indicating that immediate operator response is required.

## 6.12   Attention for operation

A. Please check the device before using, and confirm that it can work normally.

B. The finger should be in a proper position (see the attached illustration of figure 5 for reference), or else it may result in inaccurate measure.

C. The SpO2 sensor and photoelectric receiving tube should be arranged in a way with the subject’s arteriole in a position there between.

D. The SpO2 sensor should not be used at a location or limb tied with arterial canal or blood pressure cuff or receiving intravenous injection.

E. Do not fix the SpO2 sensor with adhesive or else it may result in venous pulsation and inaccurate measure of SpO2 and pulse rate.

F. Excessive ambient light may affect the measuring result. It includes fluorescent lamp, dual ruby light, infrared heater, direct sunlight and etc.

G. Strenuous action of the subject or extreme electrosurgical interference may also affect the accuracy.

H. Testee can not use enamel or other makeup.

I.  Please clean and disinfect the device after operating according to the User Manual.

## 6.13 Clinical restrictions

A. As the measure is taken on the basis of arteriole pulse, substantial pulsating blood flow of subject is required. For a subject with weak pulse due to shock, low ambient/body temperature, major bleeding, or use of vascular contracting drug, the SpO2 waveform (PLETH) will decrease. In this case, the measurement will be more sensitive to interference.

B. For those with a substantial amount of staining dilution drug (such as methylene blue, indigo green and acid indigo blue), or carbon monoxide hemoglobin (COHb), or methionine (Me+Hb) or thiosalicylic hemoglobin, and some with icterus problem, the SpO2 determination by this monitor may be inaccurate.

C. The drugs like dopamine, procaine, prilocaine, lidocaine and butacaine may also be a major factor blamed for serious error of SpO2 measure.

D. As the SpO2 value serves as a reference value for judgement of anemic anoxia and toxic anoxia, some patients with serious anemia may also report good SpO2 measurement.

# 7 Maintenance、transportation and storage

## 7.1 Cleaning and Disinfecting

Using medical alcohol to disinfect the device, nature dry or clean it with clean soft cloth.

## 7.2 Maintenance

A. Please clean and disinfect the device before using according to the User Manual(7.1).

B. Please recharge the battery when the screen shows .

C. Recharge the battery soon after the over-discharge. The device should be recharged every six months when it is no regular used. It can extend the battery life following this guidance.

D. The device needs to be calibrated once a year (or according to the calibrating program of hospital). It also can be performed at the state-appointed agent or just contact us for calibration.

## 7.3 Transportation and storage

A. The packed device can be transported by ordinary conveyance or according to transport contract. The device can not be transported mixed with toxic, harmful, corrosive material .

B. The packed device should be stored in room with no corrosive gases and good ventilation. Temperature: -40°C~60°C; Humidity: ≤95%

# 8 Troubleshooting

| Trouble | Possible Reason | Solution |
| --- | --- | --- |
| The SpO2 and Pulse Rate can not be displayed normally | The finger is not properly positioned. 2. The patient’s SpO2 is too low to be detected. | 1. Place the finger properly and try again. 2. Try again; Go to a hospital for a diagnosis if you are sure the device works all right. |
| The SpO2 and Pulse Rate are not displayed stably | 1. The finger is not placed inside deep enough. 2. The finger is shaking or the patient is moving. | 1. Place the finger properly and try again. 2. Let the patient keep calm |
| The device can not be turned on | 1. The batteries are drained or almost drained. 2. The device’s malfunction | 1.Please recharge the battery 2. Please contact the local service center. |
| The display is off suddenly | 1. The device’s malfunction. 2. The battery is drained away or almost drained away. | 1. Please contact the local service center. 2.Please recharge the battery |
| The device can not be used for full time after chagre | 1、The battery is not full charged. 2、The battery is broken | 1、Please recharge the battery 2、Please contact the local service center. |
| The battery can not be full charged even after 10 hours charging time. | The battery is broken | Please contact the local service center. |

# 9 Key of Symbols

| Signal | Description |
| --- | --- |
|  | Warning – See User Manual |
|  | The pulse oxygen saturation(%) |
|  | Pulse rate (bpm) |
|  | Full-voltage |
|  | Low-voltage |
|  | Close the alarm sound indication |
|  | Pause the alarm sound indication |
|  | Open the alarm sound indication |
|  | Close the pulse sound indication |
|  | Open the pulse sound indication |
|  | menu button/power button/function button |
|  | Type BF |
| SN | Serial number |
|  | 1. the finger clip falls off ( no finger inserted)] 2. Probe error 3. Signal inadequacy indicator |
| IPX1 | Ingress of liquids rank |
|  | WEEE (2002/96/EC) |

# Function Specification

| Information | Display Mode |
| --- | --- |
| The Pulse Oxygen Saturation（SpO2） | 2-digit digital OLED display |
| Pulse Rate（PR） | 3-digit digital OLED display |
| Pulse Intensity (bar-graph) | bar-graph OLED display |
| SpO2 Parameter Specification |
| Measuring range | 0%～100%, (the resolution is 1%). |
| Accuracy | 70%～100%：±2% ,Below 70% unspecified. |
| Average value | Calculate the Average value in every 4 measure value. The deviation between average value and true value does not exceed 1%. |
| Pulse Parameter Specification |
| Measuring range | 30bpm～250bpm, (the resolution is 1bpm) |
| Accuracy | ±2bpm or±2% (select larger) |
| Average pulse rate | Moving calculate the Average pulse rate every 4 cardio-beat,s cycle. The deviation between average value and true value does not exceed 1% |
| Safety Type | Interior Battery，ＢＦType |
| Pulse Intensity |
| Range | Continuous bar-graph display, the higher display indicate the stronger pulse. |
| Battery Requirement |
| Voltage 3.7 rechargeable lithium battery × 1 |
| Battery working life |
| Charge and discharge no less than 500 times. |
| Power Adapter |
| Input Voltage | 100 to 240 VAC, 50/60 Hz |
| Output voltage | 5 VDC |
| Output current | 250mA |
| Output power | 1.25 W |
| Oximeter Probe |
| Wavelength:660nm 880nm |
| Dimensions and Weight |
| Dimensions | 61(L) × 56(W) × 24 (H) mm |
| Weight | About 50g (with the lithium battery*1) |

# Appendix 1

| State | Alarm condition delay | Alarm signal generation delay |
| --- | --- | --- |
| Low voltage alarm | 0.6s | 20ms |
| Spo2 alarm | 400ms | 20ms |
| Pulse rate alarm | 400ms | 20ms |
| Probe error alarm | 400ms | 20ms |

Due to the constant improvement in hardware and software, actual unit may varies in looks and functionalities, please contact our customer service for update.

For further information or assistance with this product, please contact:

Simpro

Customer Service: 1-877-299-0356

Addr: 2018 Primrose Dr, Irving, TX 75063

E-mail: support@naturespiritproduct.com

Website: www.naturespiritproduct.com